## Consolidated Database
ICU/HC/CC Bed, Equipment and Supplier Database
| WCO with WHO | WCO | Equipment required (Surge List, WHO list and WCO list) | WHO | National Health Commodities Catalogue
# | WCO Categorisation | Customs Organisation Code | Item required | Incl. in WHO list? | Functional Specifications | Performance specifications
40 | Section II. Protective garments and the like | 9020 | Shields, X-Ray, Eye | No
41 | Section II. Protective garments and the like | 4015.11 | Shields, X-Ray, Glove | No
42 | Section II. Protective garments and the like | 3926.6 | Shields, X-Ray, Gonadal | No
43 | Section II. Protective garments and the like | 3926.6 | Shields, X-Ray, Throat | No
44 | Section II. Protective garments and the like | 3926.6 | Shields, Apron, X-Ray | No
45 | Section II. Protective garments and the like | 9022.14 | Shields, X-Ray, Apron | No
327 | Section II. Protective garments and the like | 4015.11 | Gloves, examination, nitrile, powder-free, pair-packed, sterile, single use (various sizes S/M/L) | Yes
325 | Section III. Disinfectants and sterilisation products | 8419.20 | Antiseptic Wipe with Alcohol & Chlorhexidine | Yes
6 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9026.80 | Oxygen flow meter, single | Yes | A medical gas regulator incorporating a flow meter. To incorporate a safety pressure relief valve. Regulator body to be manufactured from non-corrosive metal. The flow meter and flow adjustment knob to be manufactured from an impact resistant plastic (polycarbonate). High visibility flow indicator. Bullnose fitting. Nipple for direct connection of oxygen masks or tubing. | Flow meter to be graduated from 1 to 15l per minute. Must be cleanable with standard cleaning and disinfectant solutions. Expected service life: 15 years.
9 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.20 | PEEP Valves | No | Clear plastic. Autoclaveable. | Pressure adjustable from 5-20c. of H2O. Must be cleanable with standard cleaning and disinfectant solutions. And autoclaveable.
19 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.20 | Ventilator, adult, complete with humidifier and patient circuit | Yes
20 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.20 | Constant positive airways pressure (CPAP) facilities with air oxygen mixers : | Yes
21 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.20 | Ventilator,paediatric , complete with humidifier and patient circuit | Yes
32 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.20 | Nebulizers, Ultrasonic | No | General specifications: The unit must be a portable, lightweight unit of approximately 2-3kg. Ultrasonic. Should be able to connect directly to a mask. Unit to be cleanable by either boiling (5-10mins) or with a chemical disinfectant or sterilisable in autoclave at 121°C (1,2bar). Unit to include power switch to activate blower and transducer, flow rate adjustment controller. Unit to have external air inlet with replaceable high efficiency filter (50/70um). Unit to have a low operating sound level (max 60dB). Time-out function 0-60 minutes. Power requirements Mains voltage: 220V, 50-60Hz. The unit to be supplied complete with 3m power cord. | Unit to operate with maximum pressure of 1,6 bar approximately. Unit to have a nebulisation capacity of 2-5 ml per minute. Medicine chamber capacity 50-100ml. Particle size 3-8 micron. Must be cleanable with standard cleaning and disinfectant solutions. State expected Mean Time Before Failure of the unit. Expected service life: 10 years.
78 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9018.39 | Endotracheal tube introducer (adult and paeds) | Yes | Size: Adult -To fit Endotracheal tube size 6.5-11.0mm. Length:340mm-365mm. Manufactured from low friction atraumatic outer coating around a malleable aluminium core. Soft atraumatic tip, U-bend or other mechanism at proximal end to limit depth. Latex free. Clinically clean packaging. Individually packed in peel pouch.
128 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9018.39 | Oropharyngeal airways (Guedel) (various sizes 0, 1, 2, 3, 4, 5) | Yes
138 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.2 | Laryngeal masks (supraglottic airways) size 3 OR 4 OR 5 | Yes
141 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.2 | Manual bag valve mask/ manual resuscitator OR self-inflating bag with compatible masks for adults and paeds | Yes
144 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9018.39 | Simple face mask OR reservoir mask OR nasal cannula (prongs) for oxygen, (adults and paed) | Yes
211 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9018.19 | Pulse oximeter with adult & paediatric probes | Yes | Portable battery-operated hand-held robust pulse oximeter. Visual display must be brightly illuminated for low-light viewing. Display O2 saturation, pulse rate, plethysmographic wave form. | Oxygen saturation range 0-100%. Pulse range 20-300 beats per minute. Accuracy. Operating temperature 0-50°C. Must be cleanable with standard cleaning and disinfectant solutions. Expected service life: 5 years.
244 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9026.8 | Flowmeter for oxygen therapy | Yes | A medical gas regulator incorporating a flow meter. To incorporate a safety pressure relief valve. Regulator body to be manufactured from non-corrosive metal. The flow meter and flow adjustment knob to be manufactured from an impact resistant plastic (polycarbonate). High visibility flow indicator. Bullnose fitting. Nipple for direct connection of oxygen masks or tubing. | Flow meter to be graduated from 1 to 15l per minute. Must be cleanable with standard cleaning and disinfectant solutions. Expected service life: 15 years.
295 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.20 | Concentrator O2, 10 L, with accessories | Yes
297 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.20 | BiPAP, with tubing and patient interfaces for adult and paediatric, with accessories | Yes
298 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9018.90 | High Flow Nasal Cannula, with accessories | Yes
305 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.20 | Bubble humidifier | Yes
306 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.20 | Connector, biconical, symmetric, ext. diam. 7-11 mm | Yes
308 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.20 | Flow splitter, 5 flowmeters 0-2 L/min, for paediatric use | Yes
309 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.20 | Filter, heat and moisture exchanger (HMEF), high efficiency, with connectors, for adult and paediatric, single use | Yes
312 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.2 | Catheter, nasal, 40 cm, with lateral eyes, sterile, single use; different sizes: 10 Fr, 12 Fr, 14 Fr, 16 Fr, 18 Fr | Yes
313 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.2 | Venturi Mask, with percent O2 Lock and tubing , adult and paediatric and paediatric | Yes
315 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9019.20 | Airway, nasopharyngeal, sterile, single use, set with sizes of: 20 Fr, 22 Fr, 24 Fr, 26 Fr, 28 Fr, 30 Fr, 32 Fr, 34 Fr, 36 Fr | Yes
22 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9032.89 | Regulators, High Pressure Gas | Yes
146 | Section IV. Oxygen Therapy equipment and pulse oximeters | 9020.00 | Face mask for nebuliser OR face mask with nebuliser chamber for adult OR spacers for MDIs (adults and paediatrics) | Yes | None | None
296 | Section IV. Oxygen Therapy equipment and pulse oximeters | 7311.00 or 7613.00 | Medical gas cylinder, portable, for oxygen, fitted with a valve and a pressure and flow regulator | Yes | A medical gas regulator incorporating a flow meter. To incorporate a safety pressure relief valve. Regulator body to be manufactured from non-corrosive metal. The flow meter and flow adjustment knob to be manufactured from an impact resistant plastic (polycarbonate). High visibility flow indicator. Bullnose fitting. Nipple for direct connection of oxygen masks or tubing. | Flow meter to be graduated from 1 to 15l per minute. Must be cleanable with standard cleaning and disinfectant solutions. Expected service life: 15 years.
1 | Section V. Other medical devices and equipment | 9018.19 | Monitor, patient with with ECG, Resp., NIBP, IBP(x2), SaO2, Temperature, Capnography, Spirometry | Yes | The unit will automatically acquire and display vital measurements once switched on and connected to the patient. Data acquisition will continue at pre-set intervals, until switched off. A transport-type unit is envisaged. Given the operational environment, the smallest number of controls is preferred. Although not mandatory, external communication port would be an advantage – could be optional (USB or RJ). Power-on self-test. The interval between measurements must be selectable from 1 – 120 minutes. Must be compatible with Automatic External Defibrillator. Unit should have optional module to measure temperature, or alternative model incorporating temperature measurement should be proposed. Temperature in degrees Centigrade, state range. Must be suitable for adults, children and neonates. Reusable cuffs must be supplied: Large adult, adult, paediatric, neonatal. Cuffs and cuff covers must be washable, disinfectable. Minimum cuff tube length 1,5m. The unit must use the oscillometric method of measurement. Modes: Manual/stat and automatic. SpO2 sensor: Sensor: Dual LED, Infra-red/visible red, which does not heat the point of application to the patient. The following reusable sensors with cable and plug must be provided: Finger sensor for adults Finger sensor for children Earlobe sensor for adults Earlobe sensor for children The sensor lead must be approximately 3 metres long. LCD display: The display must offer clear viewing under all lighting conditions. The following must be displayed: Systolic, diastolic and mean arterial pressures (numeric) Oxygen saturation (SpO2 %) | State whether measurements are performed on inflation or deflation part of cycle. State BP measurement algorithm. Pressure measurements: Adult: Systolic: 30 – 270mm Hg, Diastolic: 20-240mm Hg, MAP: 20 – 250mm Hg Paediatric: Systolic: 30 – 160mm Hg, Diastolic: 20-120, MAP: 20 – 120mm Hg Neonate: Systolic: 30 – 130mm Hg, Diastolic: 20-100mm Hg, MAP: 20 – 120mm Hg Heart rate: 30 – 240 min. Indication of pulse irregularity desirable. Unit should have optional module to measure temperature, or alternative model incorporating temperature measurement should be proposed. Temperature in degrees Centigrade, state range. Pulse oximeter. State SpO2 measurement method. Suitable for adult, paediatric and neonatal use. The unit must continuously measure and display SpO2 % . Oxygen saturation (SpO2 %): Range of measurement: 0-100%. The accuracy must be better than 2,5 % (70-100%) and 5% (50-70%). Pulse rate (beats/min): Range of measurement: 20-300 bpm. The pulse rate accuracy must be better than ± 2 %. An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation. Automatic rejection of motion artefacts. Must be protected against electrosurgical and defibrillation voltages. The unit must have trending for at least 72 hours. Expected service life: 10 years.
2 | Section V. Other medical devices and equipment | 9018.90 | Blood warmer | No
3 | Section V. Other medical devices and equipment | 9801.00 | Examination/procedure light | No | Mobile examination light for general medical examination. LED light source. Easy switching method. Self-balancing mechanism (spring- or friction clutch-based), so that the light head holds its position wherever directed. To have sufficient swing radius to cover both ends of standard examination couch, from a standing point on floor at mid-length of the couch. Unit mounted on tilt-resistant base, with castors. Epoxy-coated steel construction. Power cord to be fixed to light, 3meter length with moulded SABS approved 15A 3-pin plug. | Daylight colour temperature of light source (6000 K). Light intensity at least 35 000 lux at 60cm. Must be cleanable and disinfectable with standard solutions.
4 | Section V. Other medical devices and equipment | 8413.19 or 9018.90 | Infusion pump, multichannel | Yes
7 | Section V. Other medical devices and equipment | 9018.90 | Glucotest machine or equivalent : | No | Portable, handheld blood glucose monitor/meter with proven accuracy. Determination of glucose in fresh capillary blood, serum, venous blood and arterial blood by reflectance photometry or amperometric enzyme electrodes, making use of electronic circuitry. Visual back-up (comparison colour in the control window on the back of the strip with colour range on strip vial label). No memory loss during battery changes. Memory capacity- minimum of 500 measurements. Neonatal Clearance from FDA for test measurements. Weight: approximately 50grams (without batteries). Display: LCD with reasonable size display and visible under different lighting conditions. Battery operated (commonly available batteries). Must use replaceable dry cell batteries. Battery indicator warning for low battery. Operator must be able to replace batteries easily without using tools. On/off switch. Button for searching through memory (blood values). Insertion slot to insert test strips. Test strip guide with an indicator that the machine is ready to perform measurements. Automatic measuring of sample. Guarantee on shelf life of test strips irrespective of whether the pack of strips is opened or not. | Home-use single user devices are not suitable. The device must be disinfectable to prevent blood-borne disease transmission. Specify validated disinfection procedures. Accuracy: 95% of all readings must be within max 15% of lab measurement (12% preferred). 99% of measurements must be within 20% of lab measurement (15% preferred). Supply studies. Measuring range 0.6 -30.0 mmol per litre. Measuring time approximately 5 seconds. Automatic cut-off after ± 120 seconds.
8 | Section V. Other medical devices and equipment | 8419.90 | Exhaled CO2 and O2 analyzers | No
10 | Section V. Other medical devices and equipment | 8413.19 | Pump, enteral feeding | Yes | No technical specifications
15 | Section V. Other medical devices and equipment | 9018.90 | Stethoscope | No | Stethoscope, binaural twin-head. Head to be one-piece stainless steel or medical-grade aluminium, with air-tight seal between head and shaft. Non-removable bell and non-corrugated diaphragm. Bell to be of shallow type, with replaceable non-chill sleeve. Diaphragm to be of good quality reinforced plastic. Rotation mechanism must have positive click-stop in both bell and diaphragm position, such that head does not rotate during auscultation. Tubing must be of sufficient thickness to attenuate environmental noise. Latex-free material. Single tube (although may be double-lumen). Longer tube length preferred (minimum 70cm from earpiece to head). Ear tubes must be adjustable (i.e. rotatable) or angled forward at 15 degrees. Tubes must be of stainless steel with screw-on earpieces. Junction between tubes and earpieces must be airtight. Headpiece spring may be internal (preferred) or external, stainless steel or chrome-plated spring steel. Earpieces should be soft, formfitting, but standard hard plastic earpieces must also be supplied. | Head must amplify low frequency <100hz by approximately 15db, attenuate higher frequencies >500hz. Performance certificate required. Entire unit must be disinfectable with standard solutions. Expected service life: 15 years.
16 | Section V. Other medical devices and equipment | 9018.90 | Suction unit, single bottle/disposable bag, wall outlet | Yes | Fixed installation. Serve number of surgeries: 2-4. Output power: Single-phase, Approximately 1.7kW-9A. Maximum flow: No less than 820 l/min. Operating head for continuous service: Maximum 3300mm H2O. Average water consumption: Approximately 6 l/min. Piping pressure: Approximately 4 bar. Dimensions: Approximately 370mm W, 300mm D, 420mm. | Maximum flow: No less than 820 l/min. Operating head for continuous service: Maximum 3300mm H2O. Expected service life 15 years.
17 | Section V. Other medical devices and equipment | 9018.90 | Transducer holder | No
18 | Section V. Other medical devices and equipment | 8413.19 | Syringe pump
23 | Section V. Other medical devices and equipment | 9018.90 | Forced air convective warming device: | No
24 | Section V. Other medical devices and equipment | 9018.19 | Ophthalmoscope and ENT diagnostic sets | No | A wall-mounted unit, combining a wall-mounted cradle, an Otoscope, an ophthalmoscope, and the necessary transformer. The wall unit must be of a durable, easy-to-clean material, such as aluminium or high-impact ABS. The unit must have cradles for the other two instruments, and optionally have a mounting point for the aneroid sphygmomanometer. The cradles should have sensors or switches to automatically turn the instruments on when lifted out of their cradles, and switch off when they are placed back in the cradles. Connecting cables between wall unit and instruments can be coiled or straight, but must have positive-lock connections to the instrument heads. Cables must be covered in latex-free rubber or PVC, no woven covering. Handle based Rheostats for light adjustments for both instrument. Cassette for specula not required. The transformer may be integrated into the wall mounting plate, or be external, for direct insertion into a wall socket. In either case, mains connecting cable minimum length 2 meters must be supplied, with a moulded plug to fit standard SABS-approved 15-amp three pin outlets. The transformer must accept 220 volts AC at 50 Hz. Guarantee on all cables and power supplies must be 5 years or better. Otoscope head to fit onto the handle of the wall mounted transformer. The Otoscope to be supplied with four polypropylene specula sizes 2.5, 3, 4 and 5mm (6 of each size, re-usable quality). The Otoscope head is to be supplied with a positive-lock connection to the handle. The Otoscope is to utilize fibre optic cool-light supply with no reflections. The Otoscope is to have a Wideangled viewing lens for instrumentation whilst under magnification. The unit is to be completely sealed for Pneumatic otoscopy (i.e. not open operating model). Light LED (optimized for best tissue viewing). Standard Ophthalmoscope head to fit onto the second handle of the wall mounted transformer. The Ophthalmoscope head is to cover -25 to 40 Diopters. Illuminated lens dial with clear indication of settings. Red-free filter. Very small, small, large, target and slit apertures. Sealed optics for dust protection. The Ophthalmoscope head is to be supplied with a positive-lock connection to the handle. Light supply LED (optimized for best tissue viewing). | Must be cleanable and disinfectable with standard solutions.
25 | Section V. Other medical devices and equipment | 3822.00 | Urine testing apparatus : | No
26 | Section V. Other medical devices and equipment | 9018.90 | A flexible fibre-optic bronchoscope : | No
27 | Section V. Other medical devices and equipment | 9018.90 | Chest drainage suction apparatus | No
28 | Section V. Other medical devices and equipment | 9018.90 | Micro-haematocrit | No
29 | Section V. Other medical devices and equipment | 9031.80 | Spirit levels | No
30 | Section V. Other medical devices and equipment | 9018.90 | Respirometers | No
31 | Section V. Other medical devices and equipment | 9018.19 | Transport monitor | Yes | The unit will automatically acquire and display vital measurements once switched on and connected to the patient. Data acquisition will continue at pre-set intervals, until switched off. A transport-type unit is envisaged. Given the operational environment, the smallest number of controls is preferred. Although not mandatory, external communication port would be an advantage – could be optional (USB or RJ). Power-on self-test. The interval between measurements must be selectable from 1 – 120 minutes. Must be compatible with Automatic External Defibrillator. Unit should have optional module to measure temperature, or alternative model incorporating temperature measurement should be proposed. Temperature in degrees Centigrade, state range. Must be suitable for adults, children and neonates. Reusable cuffs must be supplied: Large adult, adult, paediatric, neonatal. Cuffs and cuff covers must be washable, disinfectable. Minimum cuff tube length 1,5m. The unit must use the oscillometric method of measurement. Modes: Manual/stat and automatic. SpO2 sensor: Sensor: Dual LED, Infra-red/visible red, which does not heat the point of application to the patient. The following reusable sensors with cable and plug must be provided: Finger sensor for adults Finger sensor for children Earlobe sensor for adults Earlobe sensor for children The sensor lead must be approximately 3 metres long. LCD display: The display must offer clear viewing under all lighting conditions. The following must be displayed: Systolic, diastolic and mean arterial pressures (numeric) Oxygen saturation (SpO2 %) | State whether measurements are performed on inflation or deflation part of cycle. State BP measurement algorithm. Pressure measurements: Adult: Systolic: 30 – 270mm Hg, Diastolic: 20-240mm Hg, MAP: 20 – 250mm Hg Paediatric: Systolic: 30 – 160mm Hg, Diastolic: 20-120, MAP: 20 – 120mm Hg Neonate: Systolic: 30 – 130mm Hg, Diastolic: 20-100mm Hg, MAP: 20 – 120mm Hg Heart rate: 30 – 240 min. Indication of pulse irregularity desirable. Unit should have optional module to measure temperature, or alternative model incorporating temperature measurement should be proposed. Temperature in degrees Centigrade, state range. Pulse oximeter. State SpO2 measurement method. Suitable for adult, paediatric and neonatal use. The unit must continuously measure and display SpO2 % . Oxygen saturation (SpO2 %): Range of measurement: 0-100%. The accuracy must be better than 2,5 % (70-100%) and 5% (50-70%). Pulse rate (beats/min): Range of measurement: 20-300 bpm. The pulse rate accuracy must be better than ± 2 %. An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation. Automatic rejection of motion artefacts. Must be protected against electrosurgical and defibrillation voltages. The unit must have trending for at least 72 hours. Expected service life: 10 years.
35 | Section V. Other medical devices and equipment | 9018.90 | Aspirators, (Thoracic or Emergency) | No
37 | Section V. Other medical devices and equipment |  | Plasmapheresis | No
38 | Section V. Other medical devices and equipment | 9018.90 | Renal replacement therapy/haemoperfusion | No
55 | Section V. Other medical devices and equipment | 4818.90 | Urinals | No
56 | Section V. Other medical devices and equipment | 3926.90 | Bedpans | No | For urine collection from female patients. Standard Oval or Triangular contour bedpan. Stainless steel, one-piece body. Polished, smooth contours, no sharp edges. | Must be sufficiently robust to support 180kg weight across the top of the unit (while resting on a solid surface), without distorting. Must be autoclaveable.
80 | Section V. Other medical devices and equipment | 9018.9 | Magill forceps | Yes | Forceps, endotracheal, Magill`s. Serrated, fenestrated jaws, screw joint. Stainless steel with matte finish. Length: 240mm. Autoclaveable, Specification number: 1691. | 15-year service life.
82 | Section V. Other medical devices and equipment | 9018.19 | Laryngoscope, including selection of adult and paediatric layngoscope blades | Yes | Laryngoscope set handle: Handle fabricated from stainless steel. The bulb is housed in the handle head. The lamp to be of LED type. Lamp to switch on when blade is extended into operating position. Quick-release head, Batteries to be housed in handle. May have second smaller handle for paediatric blades. Spare set of batteries to be supplied. Blades to be supplied: The blade sizes 0, 1, straight fibre optic (Miller), and 1,2,3,4,5 curved fibre-optic(Macintosh pattern). Blades are to be manufactured from medical grade stainless steel. With hook-on fitting and positive-locking engagement to keep blade attached to handle in closed or operating position. Integrated, non-removable light path. | Cool white or daylight LED colour temperature. Handle must have non-slip surface or knurling. Expected service life: 15 years.
86 | Section V. Other medical devices and equipment | 9018.90 | Mosquito forceps (protected jaws to clamp pilot tube) | No | Forceps, artery, Halsted-mosquito. Curved, box joint. Fine tip. Stainless steel. Length: 125mm. Autoclaveable. | 15-year service life.
88 | Section V. Other medical devices and equipment | 9018.90 | Tracheostomy dilator | No
89 | Section V. Other medical devices and equipment | 9018.90 | Small Langenbeck retractors x 2 | No
125 | Section V. Other medical devices and equipment | 9018.90 | Patella hammer | No | Queen’s Square Pattern, Babinski pattern also acceptable. Head is disc-shaped, stainless steel or chromeplated brass, with doughnut-shaped rubber bumper. Non-hardening rubber bumper. Handle is white nylon, threaded at one end to screw into the head, pointed at the other end to test for the Babinski reflex. Handle length approximately 45cm. Nylon must be flexible, guaranteed non-hardening and shatterproof for 10 years. | Must be cleanable and disinfectable with standard cleaning solutions.
172 | Section V. Other medical devices and equipment |  | Admin set 20 drops/ml 1.8m /pack or paeds 60 drops/ml 1.8m /pack (paed) | No
174 | Section V. Other medical devices and equipment | 9018.19 | Automatic External Defibrillator (AED) OR ECG monitor and defibrillator | No | The unit being offered must be a button operation Biphasic semi-automated external defibrillator and it must be suitable for ground mobile, ambulance use and ward use in hospitals. It must be robustly constructed. It must incorporate the latest and reliable Biphasic technology. The unit offered must be able to perform impedance compensation before delivering a shock to a patient. The supplier must state the minimum and maximum defibrillator energy charge that is available on the unit being offered for when unit is in manual mode. The unit must provide clearly visible screen text messages and clear audio voice prompts in the english language to guide the user through the whole operational of the unit. The visible screen text messages must be clearly visible under all lighting conditions. The supplier must state the total weight of the unit including the battery. The unit offered must function off a reliable rechargeable battery and must include the charger with a reconditioning button in the final price. Must be provided with additional set of backup batteries to ensure continuous operation. The unit must offer automated self-checks and manual check must be carried out every time the machine is switched on as well as weekly and monthly, and all the data of the self-check must be stored in the internal memory. The weekly and monthly automated self-check must include at least the following: battery, internal electronic circuitry and energy charges. The unit must be fitted with an audible alarm if any of self-tests or battery fails as well as a LCD screen that will indicate a failure of the device. The unit must have a minimum of 10 audible and 10 visual screen prompts. The length of the pad connecting cable must be at least 0.8 meters. The unit must be supplied with a patient cable that will be able to be connected to ECG electrodes for monitoring without connecting the pads. The ECG Waveform screen must be a backlit LCD screen with a size of at least 6cm x 4cm or 6cm diagonally. The unit must be supplied with a screen protector and or protective device. The unit must be able to be configured for paediatric patients. Paediatric adaptor, pads/all paddles should be included. The unit must have a manual override with password protection. The unit must be able to download information to a computer for legal purposes. The supplier must supply a set of disposable defibrillator pads, which includes the connecting cable and the cost of this set of defibrillator pads must be included in the final bid price. The disposable defibrillator pads for the unit offered must have a shelf life of at least 2 years. A carry case constructed of durable waterproof material, which has should straps must be supplied and included in the final bid price. The carry case offered must adequately protect the unit when used by ambulance and emergency personnel in poor and wet weather conditions. | The unit offered must be lightweight less than 4kg. The rechargeable batteries must have a capacity of a minimum of 50 x 150J shocks or at least 2 hours of monitoring time. The rechargeable batteries must have a charge time of less than 3 hours. The unit must be able to withstand droppage of at least 1m on any corner. Supplier must provide details. To be splash proof. The unit must be equipped with internal memory of a minimum of 1GB. The unit must be equipped with a serial port, USB or data card system for download purposes. The software to download must be included in the bid price and supplied with every AED delivered by the successful supplier. Expected service life: at least 8 years.
175 | Section V. Other medical devices and equipment | 9018.19 | Defibrillator pads for AED OR Electrodes for ECG monitor and defibrillator | No
178 | Section V. Other medical devices and equipment | 9018.90 | Haemoglobin meter | No | The meter must be lightweight (approximately 300 grams), robust and simple to operate. Must use either replaceable dry cell batteries or rechargeable batteries. Specify the type and number of dry cell batteries, and expected operating time per charge. If rechargeable, state expected service life of batteries. If the unit has rechargeable batteries, a battery charger, operating from 220V mains supply must be provided. Cable for the battery charger must be at least two metres long and be fitted with a SABS approved 15A plug. Automatic calibration. The display must provide indication of a low battery condition. The unit must display blood haemoglobin test results in g/dl or mmol/l . State whether the unit on offer makes use of disposable cuvettes, or test strips (specify type). The meter must perform self-test/ self-calibrating on power-on. State method of self-calibration, and whether periodic user calibration is required. If so, detail method and materials required and cost thereof. The meter must automatically compensate for turbid blood samples and thereby still provide accurate measurement results. | The minimum measuring range of the meter must be as follows: 0 – 25,6g/dl 0 – 15,9mmol/l. The meter must display test results within 60 seconds. The accuracy of the meter must be ± 1,5 %.
180 | Section V. Other medical devices and equipment | 9018.90 | Blood glucometer testing strips | No
182 | Section V. Other medical devices and equipment | 9018.90 | Rescue scissors (to cut clothing) | No | Heavy duty rescue scissors suitable for Emergency Medical Care and Rescue of patients in the prehospital environment. The scissors must not be smaller or larger than 15cm long. The handle must be ergonomic to allow for a solid grip. The blades must be constructed from high quality 420 surgical grade stainless steel. Autoclaveable. Blades and handles to have non-stick surface. | The blades must be able to cut most materials including clothing and security belts. 10-year service life.
212 | Section V. Other medical devices and equipment | 9018.90 | Non-invasive electronic blood pressure monitoring device including paediatric, adult& large adult cuff sizes | Yes | The unit will automatically inflate and deflate the cuff once a switch is pressed. Automatic timed repetitive readings are not required, as this machine will not be used for ongoing monitoring. However, a unit which automatically performs several measurements and displays the average, would be desirable. Modes: Manual/stat (optionally average of multiple readings). Continuous monitoring not required. Given the operational environment, the smallest number of controls is preferred. Ideally, unit will automatically go to power-saving mode between measurements. Although not mandatory, external communication port would be an advantage. Could be optional (USB or RJ). Must be suitable for adults, children and neonates. The unit must use the oscillometric method of measurement. Unit should have optional module to measure temperature, or alternative model incorporating temperature measurement should be proposed. Temperature measured and displayed in degrees Centigrade, state range available. Values must be displayed numerically via LED or backlit LCD (colour or monochrome). Must be clearly visible under all lighting conditions. Display should also indicate battery- and charging status. Auto zero when unit is switched on reusable cuffs must be supplied: Large Adult, Adult, paediatric, neonatal. Cuffs and cuff covers must be washable, disinfectable. The air hose must be at least 1,5 metres long. The unit must charge from 220 V, 50 Hz mains supply. Rechargeable battery and charger incorporated in the unit. The battery must be charged while the unit is operating from the mains supply. Battery must be replaceable. Battery charger to be incorporated in the unit. The battery must automatically take over in the event of a mains power failure. The length of the mains cable must be a minimum of three metres. The battery must operate for at least eight hours with measurements every five minutes. Data storage: last reading, or last average. The unit must be impervious to liquids. The case of the unit must be robust and impact resistant. Must have mounting point for wall and pole. | State whether measurements are performed on inflation or deflation part of cycle. State BP measurement algorithm. Pressure measurements: Adult: Systolic: 30-270mm Hg, Diastolic: 20-240mm Hg, MAP: 20-250mm Hg Paediatric: Systolic: 30-160mm Hg, Diastolic: 20-120, MAP: 20-120mm Hg Neonate: Systolic: 30-130mm Hg, Diastolic: 20-100mm Hg, MAP: 20 – 120mm Hg Heart rate: 30-240 /min. Indication of pulse irregularity desirable Low battery: Alarm at least 10 minutes before battery power is exhausted Over pressure protection: Automatic deflation at 300mm Hg adult, 150mm Hg neonate Average measurement time: approximately 30s. If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated. Supplier to state Mean Time Before Failure for Unit and any parts which may have different MTBF.
245 | Section V. Other medical devices and equipment | 8413.19 | Multichannel infusion pump | Yes
246 | Section V. Other medical devices and equipment |  | Physiologic monitor | No
247 | Section V. Other medical devices and equipment | 9018.9 | Suction electric machine | No | Unit must be self-contained in use, and easily carried by hand. Must contain a rechargeable battery (non-memory forming) with built-in or external charger. The gauge must be marked in kPa and mmHg. Must have a battery charge level indicator. On-off switch must be illuminated and have a splash-proof cover. Must have suction control for accurate regulation of the vacuum. One receiver jar of 2l capacity, graduated in 100ml steps. Receiver jar must be high impact crack resistant polycarbonate. Receiver jar and lid must be removable and autoclavable. Receiver jar lid must have overflow protection. Receiver jar must operate without liners. Vacuum pump regulator must be fitted with a hydrophobic bacterial filter to prevent aspirated fluid from entering the suction pump. The suction inlets and connectors must be stainless steel. The pump must be oil free and maintenance free. The unit must operate from a 220 V AC supply. Must have a mains cable at least 3m long fitted with a SABS approved three pin 15A plug. Battery status indicator. Battery capacity: at least 1 hour operation. Battery must be sealed, maintenance free. The suction unit must be easy to service and maintain. | Vacuum pressure minimum 75mmHG and adjustable up to 150mmHG. The suction capacity must be at least 25l per minute. Working noise less than 50 dB.
250 | Section V. Other medical devices and equipment | 9018.90 | Adult, paediatric and large cuffs (3) for Baumanometer | Yes | The unit will automatically inflate and deflate the cuff once a switch is pressed. Automatic timed repetitive readings are not required, as this machine will not be used for ongoing monitoring. However, a unit which automatically performs several measurements and displays the average, would be desirable. Modes: Manual/stat (optionally average of multiple readings). Continuous monitoring not required. Given the operational environment, the smallest number of controls is preferred. Ideally, unit will automatically go to power-saving mode between measurements. Although not mandatory, external communication port would be an advantage. Could be optional (USB or RJ). Must be suitable for adults, children and neonates. The unit must use the oscillometric method of measurement. Unit should have optional module to measure temperature, or alternative model incorporating temperature measurement should be proposed. Temperature measured and displayed in degrees Centigrade, state range available. Values must be displayed numerically via LED or backlit LCD (colour or monochrome). Must be clearly visible under all lighting conditions. Display should also indicate battery- and charging status. Auto zero when unit is switched on reusable cuffs must be supplied: Large Adult, Adult, paediatric, neonatal. Cuffs and cuff covers must be washable, disinfectable. The air hose must be at least 1,5 metres long. The unit must charge from 220 V, 50 Hz mains supply. Rechargeable battery and charger incorporated in the unit. The battery must be charged while the unit is operating from the mains supply. Battery must be replaceable. Battery charger to be incorporated in the unit. The battery must automatically take over in the event of a mains power failure. The length of the mains cable must be a minimum of three metres. The battery must operate for at least eight hours with measurements every five minutes. Data storage: last reading, or last average. The unit must be impervious to liquids. The case of the unit must be robust and impact resistant. Must have mounting point for wall and pole. | State whether measurements are performed on inflation or deflation part of cycle. State BP measurement algorithm. Pressure measurements: Adult: Systolic: 30-270mm Hg, Diastolic: 20-240mm Hg, MAP: 20-250mm Hg Paediatric: Systolic: 30-160mm Hg, Diastolic: 20-120, MAP: 20-120mm Hg Neonate: Systolic: 30-130mm Hg, Diastolic: 20-100mm Hg, MAP: 20 – 120mm Hg Heart rate: 30-240 /min. Indication of pulse irregularity desirable Low battery: Alarm at least 10 minutes before battery power is exhausted Over pressure protection: Automatic deflation at 300mm Hg adult, 150mm Hg neonate Average measurement time: approximately 30s. If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated. Supplier to state Mean Time Before Failure for Unit and any parts which may have different MTBF.
252 | Section V. Other medical devices and equipment | 9026.10 | Peak flow meter | No | Simple aerodynamic, non-electrical device. One-way valve in mouth-piece to prevent inhalation cross-infection. Easy-to-read scale. Professional-grade instrument suitable for multi-patient screening. | Sterilisable. Measure flow from approximately 50 to 900 l/minute. Repeated readings must show <5% discrepancy. Accuracy must be within 10% of a calibrated spirometer.
256 | Section V. Other medical devices and equipment | 9018.90 | HB meter | No | The meter must be lightweight (approximately 300 grams), robust and simple to operate. Must use either replaceable dry cell batteries or rechargeable batteries. Specify the type and number of dry cell batteries, and expected operating time per charge. If rechargeable, state expected service life of batteries. If the unit has rechargeable batteries, a battery charger, operating from 220V mains supply must be provided. Cable for the battery charger must be at least two metres long and be fitted with a SABS approved 15A plug. Automatic calibration. The display must provide indication of a low battery condition. The unit must display blood haemoglobin test results in g/dl or mmol/l . State whether the unit on offer makes use of disposable cuvettes, or test strips (specify type). The meter must perform self-test/ self-calibrating on power-on. State method of self-calibration, and whether periodic user calibration is required. If so, detail method and materials required and cost thereof. The meter must automatically compensate for turbid blood samples and thereby still provide accurate measurement results. | The minimum measuring range of the meter must be as follows: 0 – 25,6g/dl 0 – 15,9mmol/l. The meter must display test results within 60 seconds. The accuracy of the meter must be ± 1,5 %.
258 | Section V. Other medical devices and equipment | 9018.90 | Urine specimen jars | No | Stainless Steel jar or flask, with broad base to resist tipping. Inverted cone design ideal. With or without hinged lid. | At least 600ml. Must be autoclaveable.
294 | Section V. Other medical devices and equipment | 9025.19 | Infrared thermometer | Yes
299 | Section V. Other medical devices and equipment | 9028.90 | Electronic drop counter, IV fluids | Yes
300 | Section V. Other medical devices and equipment | 9018.90 | Blood Gas Analyser, portable, with cartridges and control solutions | Yes
301 | Section V. Other medical devices and equipment | 9018.12 | Ultrasound, portable, w/linear and phased array cardiac transducer (5.0-7.5 MHz), w/trolley | Yes | Required for external monitoring of foetal heart rate (FHR), and external monitoring of uterine activity of pregnant patients. Twins monitoring is required. Must perform automatic self-check when switched on. There must a visible Foetal Heart Beat indicator which will flash with each detected heart beat. The unit must provide an audio signal, with variable volume, for every heart beat detected. The foetal heart rate must be measured by means of wide beam pulsed Doppler ultrasound, with auto correlation processing. Must be able to detect and display multiple FHR while discriminating between different foetuses. The transducer must be manufactured of high-impact plastic that is robust and watertight. The belts and buckles supplied for positioning of the transducer must allow easy angling for optimum positioning. Adjustable baseline position. A test button must be provided that produces a series of pulses for calibration which must be recorded and displayed. Display: > 16 cm (corner to corner). Displays must be clearly visible under all lighting conditions. Must allow Maternal Heart Rate (MHR) and Fetal Heart Rate(FHR) coincidence detection. The following parameters must be displayed: Foetal heart rate approximately 50 to 200 beats per minute Uterine activity: approximately 0 to 100 relative units. Foetal heart rate must update once per second when FHR is greater than 60 and once each beat when FHR is less than 60. The monitor must have the following controls: Volume Up/Down TOCO/IUP Zero Printer On/Off Printer fast forward. Must have audible and visual alarms for low/high FHR. Electronic storage data capacity. RECORDER Integrated dual channel thermal recorder. Dual heart rate recording for twins must be available. All dot print test. The strip chart paper must be easy to load. TRANSDUCERS Ultrasound and toco transducers must be supplied. All transducers must be of very robust design. Suppliers must guarantee shock resistance on the transducers e.g. will withstand a drop from one metre high onto a hard surface. Suppliers must state whether the ultrasound transducer is repairable and which parts are replaceable or whether it is a sealed unit that must be replaced when faulty. The ultrasound transducer cable must be a minimum of 2 m long. Connecting plug must be keyed and colour coded to avoid it being plugged or forced into the wrong receptacle on the unit. GENERAL The unit must operate off 220V, 50Hz power supply. All metal parts must be earthed. Mains cable at least three metres long with SABS approved 15A three prong plug. Must be fitted with rechargeable batteries. The unit must be fitted with an internal battery charger. Indicator for battery status. Must store trends for at least 8 hours. A cart/trolley must be supplied. The cart must have four castor wheels, two of them with brakes. The cart must have a basket/holder for accessories. | The oscillator frequency < 1,2 MHz. Ultrasound intensity must be less than 10 mW/cm2. Foetal heart rate: Range: approximately 50 to 210 beats per minute Vertical chart: at least 7 cm Vertical scale sensitivity: approximately 20 bpm per cm Resolution: 1 bpm. Uterine activity: Range: 0 – 100 units Vertical chart: at least 4 cm Vertical scale sensitivity: approximately 25 relative units per cm Resolution: 1 relative unit. Paper chart speed must be selectable for 1, 2 and 3 cm per minute. The time and date must be printed by the recorder at least every 10 minutes. Recording time must be at least 24 hours at 1 cm per minute. The unit must have a facility for remote event marking.
302 | Section V. Other medical devices and equipment | 9018.90 | Drill, for vascular access, with accessories adult and paediatric, with transport bag | Yes | Cordless 18Volt drill, 2-speed with hammer-drill function. Lithium-ion battery – at least 2 amp-hours. Forward/reverse switch. Trigger-operated variable speed control. Gearbox least 2-speed, slide – or leverselectable. Torque at least 15Nm. Adjustable clutch. No load speed on high: minimum 1500 rpm. Chuck 13mm. | Expected service life 10 years.
303 | Section V. Other medical devices and equipment | 8419.20 | Autoclave, 40-60L / 90L, with accessories | Yes
330 | Section V. Other medical devices and equipment | 3926.90 | Bag, collecting, urine, with outlet tap, with non-return valve, 2000 mL, adult, non-sterile, single use | Yes
331 | Section V. Other medical devices and equipment | 7324.90 | Basin kidney, stainless steel, 825 mL | Yes | Kidney dish. Dimensions approximately: Large: approx 250x130x40mm Small: approx 150x90x32mm. Stainless steel. Autoclaveable | 15-year service life.
5 | Section VI. Other Medical Consumables | 9018.39 | Endotracheal tube cuff inflator / pressure monitor | Yes | The unit will automatically inflate and deflate the cuff once a switch is pressed. Automatic timed repetitive readings are not required, as this machine will not be used for ongoing monitoring. However, a unit which automatically performs several measurements and displays the average, would be desirable. Modes: Manual/stat (optionally average of multiple readings). Continuous monitoring not required. Given the operational environment, the smallest number of controls is preferred. Ideally, unit will automatically go to power-saving mode between measurements. Although not mandatory, external communication port would be an advantage. Could be optional (USB or RJ). Must be suitable for adults, children and neonates. The unit must use the oscillometric method of measurement. Unit should have optional module to measure temperature, or alternative model incorporating temperature measurement should be proposed. Temperature measured and displayed in degrees Centigrade, state range available. Values must be displayed numerically via LED or backlit LCD (colour or monochrome). Must be clearly visible under all lighting conditions. Display should also indicate battery- and charging status. Auto zero when unit is switched on reusable cuffs must be supplied: Large Adult, Adult, paediatric, neonatal. Cuffs and cuff covers must be washable, disinfectable. The air hose must be at least 1,5 metres long. The unit must charge from 220 V, 50 Hz mains supply. Rechargeable battery and charger incorporated in the unit. The battery must be charged while the unit is operating from the mains supply. Battery must be replaceable. Battery charger to be incorporated in the unit. The battery must automatically take over in the event of a mains power failure. The length of the mains cable must be a minimum of three metres. The battery must operate for at least eight hours with measurements every five minutes. Data storage: last reading, or last average. The unit must be impervious to liquids. The case of the unit must be robust and impact resistant. Must have mounting point for wall and pole. | State whether measurements are performed on inflation or deflation part of cycle. State BP measurement algorithm. Pressure measurements: Adult: Systolic: 30-270mm Hg, Diastolic: 20-240mm Hg, MAP: 20-250mm Hg Paediatric: Systolic: 30-160mm Hg, Diastolic: 20-120, MAP: 20-120mm Hg Neonate: Systolic: 30-130mm Hg, Diastolic: 20-100mm Hg, MAP: 20 – 120mm Hg Heart rate: 30-240 /min. Indication of pulse irregularity desirable Low battery: Alarm at least 10 minutes before battery power is exhausted Over pressure protection: Automatic deflation at 300mm Hg adult, 150mm Hg neonate Average measurement time: approximately 30s. If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated. Supplier to state Mean Time Before Failure for Unit and any parts which may have different MTBF.
63 | Section VI. Other Medical Consumables | 9018.39 | Selection of endotracheal tubes (various sizes uncuffed 2.0mm to 5.5mm and cuffed 3.0mm to 8.0mm) | Yes
84 | Section VI. Other Medical Consumables |  | Macintosh or equivalent spray device | No
123 | Section VI. Other Medical Consumables | 3006.70 | Water-soluble lubricant / lubricating jelly | Yes
124 | Section VI. Other Medical Consumables | 9018.39 | Tape to hold tie endotracheal tube in place | Yes
148 | Section VI. Other Medical Consumables | 9018.39 | Intravenous cannula 18g green and appropriate strapping (18g to 24g) | No
152 | Section VI. Other Medical Consumables | 9018.31 | Syringes 3-part: (2ml to 20ml) | No
156 | Section VI. Other Medical Consumables | 9018.31 | Syringes: insulin syringes | No
157 | Section VI. Other Medical Consumables | 9018.39 | Needles: (various sizes 18 to 23) | No
162 | Section VI. Other Medical Consumables | 3006.10 | Suture chromic g0/0 or g1/0 1/2 75cm | No
163 | Section VI. Other Medical Consumables | 3006.10 | Suture nylon g2/0 or g3/0 3/8 45cm or g4/0 3/8 45cm | No
165 | Section VI. Other Medical Consumables | 9018.39 | Suction catheters: (various sizes F6 - F16) | No
185 | Section VI. Other Medical Consumables | 3005.10 | Wound care (gauze, bandages, cotton wools, plasters, alcohol swabs and antiseptic solutions) | No
198 | Section VI. Other Medical Consumables | 9018.39 | Urinary (Foley’s) catheter silicone/latex (various sizes 8f - 22f) | Yes
206 | Section VI. Other Medical Consumables | 9018.39 | Urinary bag specified in the surgical supply list | Yes
207 | Section VI. Other Medical Consumables | 9018.39 | Nasogastric tubes:(various sizes: 600mmfg8 or 1000mmfg10 or 1000mmfg12 ) | Yes
210 | Section VI. Other Medical Consumables | 9018.39 | Medication/vaculitre stickers | No
307 | Section VI. Other Medical Consumables | 9018.39 | Tubing, medical gases, int. diam. 5 mm | Yes
310 | Section VI. Other Medical Consumables | 3006.70 | Conductive gel, container | Yes
311 | Section VI. Other Medical Consumables | 9018.5 | Tubing, silicone | Yes
316 | Section VI. Other Medical Consumables | 9027.80 | Colorimetric End Tidal CO2 detector, adult and paediatric, single use | Yes
317 | Section VI. Other Medical Consumables | 9018.39 | Cricothyrotomy, set, emergency, 6 mm, sterile, single use | Yes
318 | Section VI. Other Medical Consumables | 9018.39 | Syringe, Luer slip, 10 mL, sterile, single use | Yes
319 | Section VI. Other Medical Consumables | 9018.39 | Arterial Blood Sample Kits (syringes with dry lithium heparin [23.5 iu/mL]) | Yes
320 | Section VI. Other Medical Consumables | 9018.39 | Central Venous Catheters kit with: Finder needle, syringe, wire, dilator, lidocaine, scalpel, needle, thread | Yes
321 | Section VI. Other Medical Consumables | 3005.10 | Transparent adhesive plasters, wash proof, 5x5 cm | Yes
322 | Section VI. Other Medical Consumables | 9018.39 | Syringe, feeding, LDT, ENFit, sterile, single use, sizes: 1 mL, 2.5 mL, 5 mL | Yes
323 | Section VI. Other Medical Consumables | 9018.39 | Syringe, feeding, ENFit, sterile, single use, sizes: 10 mL, 20 mL, 60 mL | Yes
324 | Section VI. Other Medical Consumables | 3005.10 | Pad, absorbent | Yes
326 | Section VI. Other Medical Consumables | 3924.90 | Drape, surgical, non woven, sterile, single use | Yes
39 | Section VII. Vehicles | 8705.9 | Radiographic Units, Mobile | No
59 | Section VII. Vehicles | 8713.10 | Wheelchairs | No | Porter chair for institutional use. Size 51cm wide. | Expected service life: 10 years.
47 | Section VIII. Other | 9402.90 | Matress cover | No | To be constructed of flexible. Overall dimensions: Length: 1 980 mm (-0 and +40mm) Width: 860 mm (-0 and +20mm) Thickness: 150 mm (-0 and +5mm). | Mattress must support a minimum patient mass of 180 kg. 200kg preferred. Must resume original shape when not in use. Mattress must be radio-translucent.
49 | Section VIII. Other |  | Pressure relieving mattress system, alternating | No
50 | Section VIII. Other | 9402.90 | Kick bucket | No | Mounted on four legs fitted with castors. Castors must comply with the latest issue of SANS 621, proof of compliance must be submitted. Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm and castors shall be fixed to the tube by one of the following methods: a. Solid plug (long) complying with SANS 621 subsection 3.5.6 or b. Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner c. Rubber or plastic expanding sleeves for fitting castors are not acceptable. The horizontal braces to be arranged in a flat cross formation. | Service life minimum 10 years. Must be cleanable and disinfectable with standard solutions.
52 | Section VIII. Other | 9403.20 | Locker, bedside | No | Steel bedside locker. Mild steel construction, with epoxy/nylon powder-coated finish. Materials which are easily cleaned and of lasting finish, and low noise generating nature, must be used. Stainless steel top, with rail on three sides. Single pull-out drawer, with durable sliding mechanism and stop to prevent total extraction of drawer. Locker mounted on 4 castors.
53 | Section VIII. Other | 9402.90 | Overbed table adjustable | No | Cantilever-type over-bed table – can be placed in position from the side of the bed. | Adjustable height from 740 mm to 1 090 mm. Automatic lock. Minimum weight bearing capacity measured at free end of table: 20kg without flexing or distortion of frame.
54 | Section VIII. Other | 9402.90 | Bed Cradles | No
57 | Section VIII. Other | 9402.90 | anaesthetist stool | No | Stool, anaesthetist, with castors. To comply with CKS233 Type B. Aluminium flat seat. Spindle to be threaded with square or acme thread. Castors: 75mm. Castors must comply with the latest issue of SANS 621, proof of compliance must be submitted. Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm and castors shall be fixed to the tube by one of the following methods: a. Solid plug (long) complying with SANS 621 subsection 3.5.6 or b. Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner c. Rubber or plastic expanding sleeves for fitting castors are not acceptable. | Height adjustments, floor to top of seat: 485mm to 685mm. Service life minimum 10 years. Must be cleanable and disinfectable with standard solutions.
58 | Section VIII. Other | 9402.90 | emergency trolley | No | Dedicated “crash cart” type trolley. All metal or ABS construction Side-accessible slide-out tray for medication preparation Top drawer lockable half-height, with Moulded insert for emergency drugs. Facility for Tamper– proof seals Four half-height drawers, for resuscitation equipment, bandages, syringes, needles, ET tubes One deep drawer for manual resuscitation ventilators and masks Side bracket for portable suction machine Side accessory tilt-out drawers On-board attached multi-plug with at least 5 15A 3-pin sockets (switchable) and 5 Euro-style sockets. Single power lead, 3-core 2,5mm2, terminated with moulded 15A 3-pin plug (SANS approved) Raised bridge rail at rear of top, to attach NIBP monitor and AED Swing-out trays not suitable Built-in extendable drip stand, left- or right-mounting Mounted on 4 approved castors, with non-loosening fixings Hooks for mounting cardiac board at rear Oxygen cylinder holder not required Cleanable and disinfectable with standard solutions. | Service life minimum 10 years. Must be cleanable and disinfectable with standard solutions. Dimensions must allow passage through standard doorway.
60 | Section VIII. Other | 3926.90 | Sharp waste disposal unit | No | Puncture-proof plastic (polypropylene), yellow with red contrast. Tamper-proof, leak-proof. Lid cannot be re-opened once closed. Spill-proof (entry port should not permit contents to fall out if the bin is knocked over). Labelled “Danger, contaminated sharps”. International hazardous waste symbol. | Cleanable with standard cleaning and disinfectant solutions. Must withstand Hot Water Wash (90°C).
62 | Section VIII. Other | 9402.90 | Medicine trolley | No | Multi-drawer mobile medicine trolley, for use in a ward environment. Flat top surface, for writing and dispensing. Must have 4-6 drawers with adjustable dividers, 2 large storage drawers without dividers. Drawers must have smooth, durable sliding mechanism with stop to prevent total extraction of drawer. Mounted on 4 heavy duty swivel castors 125 mm diameter, the two front castors should be fitted with foot brakes with toe release mechanism. Approximate dimensions: Height: 950 mm Width: 750 mm Depth: 500 mm. Materials which are easily cleaned and of lasting finish, and low noise generating nature, shall be used.
90 | Section VIII. Other | 9402 | Chair | No | Must comply with SANS 660:12, with the following additions/modifications: The chair to be manufactured from mild steel tubing, which is to be chrome or black powder-coated epoxy-finished, with replaceable plastic anti-scratch anti-noise protector clips The frame should preferably be of the sled type, anti-tilt in normal usage No arms Stackable design preferred Chair to have fully moulded shell Seat must be replaceable. | Chair must support a body weight of 200 kg without flexing or deforming. Expected service life of at least 5 years. Must be cleanable and disinfectable with standard solutions.
91 | Section VIII. Other | 8418.10 | Refrigerator for medicines | No | Upright type, single door, approximately 260l. Operating temperature: +2°C to +8°C at ambient temperature +5°C to 35°C and +60% humidity. Optionally the supplier can propose more than one model for use in moderate and hot temperature zones. Automatic defrost. Thermostat control should be graduated in degrees C, with 1 degree increments. Display accuracy: 0.1C. Ideally fridge and freezer compartments should have separate controls. Glass door not a prerequisite. Self-closing, lockable door. Forced air circulation system with intelligent control. Audible and visual alarm for high or low temperature. Door ajar alarm, system failure alarm. 4 casters mounted under the bottom (can be 2 fixed, 2 rotating). Interior LED or fluorescent lighting. 5/6units shelves made of quality plastic- or epoxy- coated steel wire. Exterior made from high quality structural plate, interior made from ABS material. Visual and audible warning device for high and low temperature alarm (Min/Max thermometer). Safety thermostat to avoid minus temperatures. Built-in elements must be resistant to mechanical damage. Good quality compressor (1 unit) and fan motor. Power supply: 220V/50Hz 1 phase. | Noise emission: < 60 dB(A). Warning device for min. 12h in the event of power failure (power failure alarm). Temperature holdover period 4 hours or better, in case of power failure. Must be cleanable and disinfectable with standard materials.
96 | Section VIII. Other |  | Soap dispenser | No | Simple, heavy-duty stainless steel cage-type unit, without hinged lever arm. Cage to be securely welded to stainless steel wall-mounting plate, with 3-4cm projection from frontsurface of plate. May be canted forwards, for easier elbow access to pump mechanism on bottle. Cage must accept non-proprietary standard, replaceable refillable, pump bottles, approximately 65mm diameter, about 500ml volume. Mounting plate to be pre-drilled for wall mounting. | Cleanable with standard cleaning and disinfectant solutions. Expected service life: 15 years.
98 | Section VIII. Other | 9402.90 | Paper towel dispenser | No | Wall-mounted polycarbonate dispenser for folded paper towels. | Cleanable with standard cleaning and disinfectant solutions. Chemical resistant. Expected service life: 15 years.
107 | Section VIII. Other |  | Spare bulbs for Laryngoscope (NA if the laryngoscope has a built in bulb) | No
108 | Section VIII. Other |  | Spare batteries for laryngoscope handle | No
183 | Section VIII. Other | 3919.90 | Paediatric Broselow tape OR PAWPER TAPE | No | Durable tape with standard length markings and colour zones. Each colour zone is associated with recommended dosages of resuscitation drugs and defibrillation energy for use in infants, where the weight is unknown. Flexible type preferred, for storage purposes. Non-smudge, non-fade. Must be cleanable with detergents and standard disinfectants. Supplier to state which cleaning solutions are unsuitable. Supplier to state whether the basis for the year edition of the tape (the US National Health and Nutrition Examination Survey – NHANES data) has been correlated with similar South African data. | Supplier to state year edition of tape. Supplier to state whether all required clinical information is printed on tape, or whether a laminated card with resuscitation algorithms is required. In the latter case, supplier to propose such a card, specifying cost.
241 | Section VIII. Other | 9018.50 | Headlight | No | Lightweight head-worn LED light. Adjustable headband, with over-the-head strap to prevent slippage. Tilt-adjustable. CREE LED element. Daylight colour temperature preferred. Easy access on-off switch on lamp. Easy access focus ring on lamp. Rechargeable Lithium-ion batteries – on headband or belt-worn. Should also accept equivalent size of alkaline cells in emergency. Intermittent flash, different power modes not required. Splashproof, at least Ingress Protection 44 rating. | Max weight of headgear less than 180gm. Minimum light output 200 lumens. Operating time at least 5 hours on full output. Focusable from narrow to wide beam. Beam must project light uniformly at 1 metre – no rings or dark spots. Must be cleanable and disinfectable.
249 | Section VIII. Other | 9018.50 | Wall mounted or portable angle poise style examination lamp | No | Wall-mounted examination light for general medical examination. LED light source. Easy switching method. Self-balancing mechanism (spring or friction clutch-based), so that the light head holds its position wherever directed. To have sufficient swing radius to cover both ends of standard examination couch, from a mounting point on wall at mid-length of the couch. | Daylight colour temperature of light source (~6000K). Light intensity at least 35 000 lux at 60cm. Power cord to be fixed to light, 2m length with moulded SABS approved 15A 3-pin plug. Must be cleanable and disinfectable with standard solutions.
253 | Section VIII. Other | 9018.50 | Tuning fork (only required in one consultation room) | No | Tuning fork to be manufactured from medical grade stainless steel. Attached flat base for vibration and bone conduction tests. Models without weights on tines preferred. Set of one each 128Hz and 256Hz . | Each fork to be clearly marked with vibration frequency. Accurate to 0,2 HZ. Expected service life: 20 years.
254 | Section VIII. Other | 9017.80 | Tape measure | No | Flat tailor-type tape measure, in durable flexible plastic or fiberglass. Durable, non-fade, non-smudge, high-contrast markings in millimetres and centimetres. The ends of the tape may be reinforced with plastic or metal strips. | Non-stretch tape. Length must be at least 150cm. Must be cleanable with standard cleaning and disinfectant solutions.
255 | Section VIII. Other | 8423.10 | Adult clinical scale up to 150 kg | No | Scale, with large dial. Comply with the specifications on Appendix H, Type 2, see attached. Capacity: at least 150kg, dial to be at least 150mm in diameter. Must include height measuring rod. Mounted on castors, roller size: 30-50mm. Fixed direction, scale to be able to tilt for moving. | Weighing capacity up to 150kg. Expected service life: 10 years. Accuracy: within 1gm per 1000gm. Repeatability: Standard deviation less than 10gm after 20 measurements at 50kg.
257 | Section VIII. Other | 9017.80 | Height measure | No | Stadiometer (Height board). To measure height of children and adults while standing. Range: 20-125cm, graduated in m, cm and mm, with numbers at metre and cm positions. Measuring surface must be free of all moving parts that may affect the markings. Movable headboard must slide easily on the vertical section and lock in the measurement position. Headboard must lock at 90° to the vertical measuring section. Locking mechanism must not affect the measuring surface. Sturdy and stable footplate. Must allow the heels, buttocks and back of the head to make contact with the vertical surface during measurement. Must be made of durable, non-flexible, smooth material that is water resistant, easy to clean, with no sharp edges or unfinished parts. | Must be accurate to at least 0,5mm per metre. Expected service life 15 years.
259 | Section VIII. Other | 8423.10 | Baby scale in paeds wards | No | Scale, weighing electronic, baby and child convertible. Battery operated. Automatic switch off. Increments: 10g up to 10kg, 20g > 10kg LCD display: digits at least 20mm high. Electronic damper. Hold feature to retain weight reading. Function to display breast milk intake. | Weighing capacity up to 20kg. Must be accurate within 10gm at 20kg. Repeatability: Standard deviation less than 10gm after 20 measurements at 10kg. Expected service life 10 years.
260 | Section VIII. Other | 9404.90 | Bed mattress | No | To be constructed of flexible. Overall dimensions: Length: 1 980 mm (-0 and +40mm) Width: 860 mm (-0 and +20mm) Thickness: 150 mm (-0 and +5mm). | Mattress must support a minimum patient mass of 180 kg. 200kg preferred. Must resume original shape when not in use. Mattress must be radio-translucent.
261 | Section VIII. Other | 9402.90 | Hospital bed | No | Mild steel frame with epoxy/nylon powder coating colours: white, cream or grey. To be fitted with castors, two swivel, two locking. Removable head and foot ends (ABS material is acceptable). Mattress support: mattress support other than weldmesh is required, provide details. | Bed must support a minimum patient mass of 180 kg. 200kg preferred. Adjustable backrest with gas spring assist, suitable for 100 kg patient. Bed must not disassemble if sides or ends are lifted – collapsible “emergency” beds are not fit for this purpose.
263 | Section VIII. Other | 9402.90 | Bedrail | No
264 | Section VIII. Other | 9402.90 | Bedside cabinet | No | Steel bedside locker. Mild steel construction, with epoxy/nylon powder-coated finish. Materials which are easily cleaned and of lasting finish, and low noise generating nature, must be used. Stainless steel top, with rail on three sides. Single pull-out drawer, with durable sliding mechanism and stop to prevent total extraction of drawer. Locker mounted on 4 castors.
265 | Section VIII. Other | 9401.79 | Footstool | No | Stool, foot, double step. To comply with the specification CKS 233 in Annexure J, see attached. Aluminium chequered plate top. Stainless steel edging. Mild steel frame and legs with epoxy/nylon powder-coated finish, to comply with SANS 778 paragraph 5.2, proof of compliance must be submitted. Heavy-duty rubber non-slip, non-scratching ferrules on feet. | Service life minimum 10 years. Must be cleanable and disinfectable with standard solutions.
266 | Section VIII. Other | 9402.10 | Intravenous stand | No | Drip stand, mobile. Mild steel frame with epoxy powder-coated or nylon powder-coated finish to comply with SANS 778 paragraph 5.2, proof of compliance (Test Report) must be submitted. Castors must comply with the latest issue of SANS 621, proof of compliance (Test Report) must be submitted. Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm and castors shall be fixed to the tube by one of the following methods: a. Solid plug (long) complying with SANS 621 subsection 3.5.6 or b. Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner c. Rubber or plastic expanding sleeves for fitting castors are not acceptable. Drag chain to be supplied. | Must be cleanable and disinfectable with standard solutions. Expected service life: 15 years.
272 | Section VIII. Other | 9402.90 | Screen for examination table | No | Four-panel concertina- foldable screen, consisting of steel frames on castors, and plastic curtains. Must comply with the latest issue of SANS 897, fig.3, Type C. Overall height 1 680 mm. Mild steel frame and legs with epoxy/nylon powder-coated finish, to comply with SANS 778 paragraph 5.2, proof of compliance must be submitted. Castors: 75 mm Castors must comply with the latest issue of SANS 621, proof of compliance must be submitted. Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm and castors shall be fixed to the tube by one of the following methods: a. Solid plug (long) complying with SANS 621 subsection 3.5.6 or b. Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner c. Rubber or plastic expanding sleeves for fitting castors are not acceptable. Curtain to be as follows: Strong non-transparent plastic Sides to be strengthened and to be full fit in frame sections. Curtain must be easily removable. | Curtains must be washable. Expected service life: Frames 10 years Curtains 3 years.
275 | Section VIII. Other | 9402.90 | Dressing trolley | No | Trolley, medical dressing. To comply with the latest issue of SANS 726 (see fig 4). Length: 915mm (± 5mm). Width: 455mm (± 5mm). The trolley is open and has two removable stainless steel shelves. Durable rubber plugs (or similar noise reducing material) must be inserted into rails below removable shelves. It has a continuous guard rail along the sides and back of the top shelf and a continuous rail fitted around the bottom, above the bottom shelf. Materials which are easily cleaned and of lasting finish, and low noise generating nature, shall be used. Stainless steel throughout. Four swivel castors with a diameter of 75mm each. Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm and castors shall be fixed to the tube by one of the following methods: a. Solid plug (long) complying with SANS 621 subsection 3.5.6 or b. Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner. | Service life minimum 10 years. Must be cleanable and disinfectable with standard solutions.
276 | Section VIII. Other | 9402.90 | Instrument cabinet | No
278 | Section VIII. Other | 9402.90 | Linen/Laundy Trolley | No | Castor-mounted, epoxy-coated mild-steel trolley. Vertical members support a horizontal circular or square frame, onto which a woven polyester soiled linen receptacle attaches. The receptacles has eyelets for attachment to the frame, as well as a draw-cord and stitched-on handles. Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm and castors shall be fixed to the tube by one of the following methods: a. Solid plug (long) complying with SANS 621 subsection 3.5.6; or b. Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner. (Rubber or plastic expanding sleeves for fitting castors are not acceptable.) | Cleanable with standard cleaning and disinfectant solutions. Expected service life: 15 years for trolley body.
279 | Section VIII. Other |  | Medication dispenser | No
280 | Section VIII. Other |  | Medication trolley | No | Multi-drawer mobile medicine trolley, for use in a ward environment. Flat top surface, for writing and dispensing. Must have 4-6 drawers with adjustable dividers, 2 large storage drawers without dividers. Drawers must have smooth, durable sliding mechanism with stop to prevent total extraction of drawer. Mounted on 4 heavy duty swivel castors 125 mm diameter, the two front castors should be fitted with foot brakes with toe release mechanism. Approximate dimensions: Height: 950 mm Width: 750 mm Depth: 500 mm. Materials which are easily cleaned and of lasting finish, and low noise generating nature, shall be used.
281 | Section VIII. Other | 9402.90 | Patient record rack | No
282 | Section VIII. Other | 8302.49 | Bedpan rack | No
285 | Section VIII. Other | 9402.90 | Bassinet for paeds wards | No | Bassinet, hospital, type A, with stand complete with mattress. To comply with the latest issue of CKS 237, type A Bassinet: durable, transparent plastic material Mild steel frame with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2, proof of compliance must be submitted Mattress: Must be in accordance with item RT 24-02-021 On castors. | Expected service life of frame: 15 years.
288 | Section VIII. Other | 8471.41 | Communication technology: Personal computer | No